Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

EESTI STANDARD	EVS-EN ISO 14971:2019

MEDITSIINISEADMED. RISKIHALDUSE RAKENDAMINE MEDITSIINISEADMETELE

Medical devices - Application of risk management to medical devices (ISO 14971:2019)

# EVS-EN ISO 14971:2019

Standardite reprodutseerimise ja levitamise õigus kuulub Eesti Standardikeskusele Andmete paljundamine, taastekitamine, kopeerimine, salvestamine elektroonsesse süsteemi või edastamine ükskõik millises vormis või millisel teel ilma Eesti Standardikeskuse kirjaliku loata on keelatud. Kui Teil on küsimusi standardite autorikaitse kohta, võtke palun ühendust Eesti Standardikeskusega: Koduleht www.evs.ee; telefon 605 5050; e-post info@evs.ee The right to reproduce and distribute standards belongs to the Estonian Centre for Standardisation No part of this publication may be reproduced or utilized in any form or by any means, electronic or mechanical, including photocopying, without a written permission from the Estonian Centre for Standardisation. If you have any questions about copyright, please contact Estonian Centre for Standardisation: Homepage www.evs.ee; phone +372 605 5050; e-mail info@evs.ee

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

EESTI STANDARDI EESSÕNA	 NATIONAL FOREWORD

| See Eesti standard EVS-EN ISO 14971:2019 sisaldab Euroopa standardi EN ISO 14971:2019 ingliskeelset teksti. | This Estonian standard EVS-EN ISO 14971:2019 consists of the English text of the European standard EN ISO 14971:2019. |
| --- | --- |
| Standard on jõustunud sellekohase teate avaldamisega EVS Teatajas. Euroopa standardimisorganisatsioonid on teinud Euroopa standardi rahvuslikele liikmetele kättesaadavaks 18.12.2019. | This standard has been endorsed with a notification published in the official bulletin of the Estonian Centre for Standardisation. Date of Availability of the European standard is 18.12.2019. |
| Standard on kättesaadav Eesti Standardikeskusest. | The standard is available from the Estonian Centre for Standardisation. |

Tagasisidet standardi sisu kohta on võimalik edastada, kasutades EVS-i veebilehel asuvat tagasiside vormi või saates e-kirja meiliaadressile standardiosakond@evs.ee.

# ICS 11.040.01

| EUROPEAN STANDARD NORME EUROPÉENNE | EN ISO 14971 |
| --- | --- |

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

December 2019

ICS 11.040.01

Supersedes EN ISO 14971:2012

English version

Medical devices - Application of risk management to

medical devices (ISO 14971:2019)

Dispositifs médicaux - Application de la gestion des

risques aux dispositifs médicaux (ISO 14971:2019)

Medizinprodukte - Anwendung des

Risikomanagements auf Medizinprodukte (ISO

14971:2019)

This European Standard was approved by CEN on 5 August 2019.

CEN and CENELEC members are bound to comply with the CEN/CENELEC Internal Regulations which stipulate the conditions for

giving this European Standard the status of a national standard without any alteration. Up-to-date lists and bibliographical

references concerning such national standards may be obtained on application to the CEN-CENELEC Management Centre or to

any CEN and CENELEC member.

This European Standard exists in three official versions (English, French, German). A version in any other language made by

translation under the responsibility of a CEN and CENELEC member into its own language and notified to the CEN-CENELEC

Management Centre has the same status as the official versions.

CEN and CENELEC members are the national standards bodies and national electrotechnical committees of Austria, Belgium,

Bulgaria, Croatia, Cyprus, Czech Republic, Denmark, Estonia, Finland, France, Germany, Greece, Hungary, Iceland, Ireland, Italy,

Latvia, Lithuania, Luxembourg, Malta, Netherlands, Norway, Poland, Portugal, Republic of North Macedonia, Romania, Serbia,

Slovakia, Slovenia, Spain, Sweden, Switzerland, Turkey and United Kingdom.

CEN-CENELEC Management Centre:

Rue de la Science 23, B-1040 Brussels

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EUROPÄISCHE NORM

© 2019 CEN/CENELEC 	All rights of exploitation in any form and by any means 	Ref. No. EN ISO 14971:2019 E reserved worldwide for CEN national Members and for CENELEC Members.

EN ISO 14971:2019 (E) EVS-EN ISO 14971:2019

## European foreword

This document (EN ISO 14971:2019) has been prepared by Technical Committee ISO/TC 210 "Quality management and corresponding general aspects for medical devices" in collaboration with Technical Committee CEN/CLC/JTC 3 “Quality management and corresponding general aspects for medical devices” the secretariat of which is held by NEN.

This European Standard shall be given the status of a national standard, either by publication of an identical text or by endorsement, at the latest by June 2020, and conflicting national standards shall be withdrawn at the latest by June 2020.

Attention is drawn to the possibility that some of the elements of this document may be the subject of patent rights. CEN shall not be held responsible for identifying any or all such patent rights.

This document supersedes EN ISO 14971:2012.

According to the CEN-CENELEC Internal Regulations, the national standards organizations of the following countries are bound to implement this European Standard: Austria, Belgium, Bulgaria, Croatia, Cyprus, Czech Republic, Denmark, Estonia, Finland, France, Germany, Greece, Hungary, Iceland, Ireland, Italy, Latvia, Lithuania, Luxembourg, Malta, Netherlands, Norway, Poland, Portugal, Republic of North Macedonia, Romania, Serbia, Slovakia, Slovenia, Spain, Sweden, Switzerland, Turkey and the United Kingdom.

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Endorsement notice

The text of ISO 14971:2019 has been approved by CEN as EN ISO 14971:2019 without any modification.

3

EVS-EN ISO 14971:2019

## Contents 	Page

Foreword ........................................................................................................................................................................................................................................iv Introduction ................................................................................................................................................................................................................................vi 1 Scope .................................................................................................................................................................................................................................1

- Normative references ......................................................................................................................................................................................1

- Terms	and	definitions .....................................................................................................................................................................................1

- General requirements for risk management system ......................................................................................................7

  - Risk management process  ........................................................................................................................................................ 7 4.2 Management responsibilities ..................................................................................................................................................... 8

  - Competence of personnel .............................................................................................................................................................. 9

  - Risk management plan .................................................................................................................................................................. 9

  - Risk management file  ................................................................................................................................................................10

- Risk analysis  ..........................................................................................................................................................................................................10

  - Risk analysis process  ..................................................................................................................................................................10

  - Intended use and reasonably foreseeable misuse  ..........................................................................................10 5.3 Identification of characteristics related to safety  ................................................................................................11

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Identification of hazards and hazardous situations  ......................................................................................11

  - Risk estimation  .................................................................................................................................................................................11 6 Risk evaluation ....................................................................................................................................................................................................12

- Risk control  .............................................................................................................................................................................................................12

7.1 Risk control option analysis ....................................................................................................................................................12 7.2 Implementation of risk control measures ..................................................................................................................13

  - Residual risk evaluation .............................................................................................................................................................13

  - Benefit-risk analysis .......................................................................................................................................................................14 7.5 Risks arising from risk control measures ...................................................................................................................14

7.6 Completeness of risk control  ................................................................................................................................................14 8 Evaluation of overall residual risk  ..................................................................................................................................................14

- Risk management review ..........................................................................................................................................................................15

- Production and post-production activities..............................................................................................................................15

10.1 General ........................................................................................................................................................................................................15 10.2 Information collection ...................................................................................................................................................................15

  - Information review ..........................................................................................................................................................................16

  - Actions .........................................................................................................................................................................................................16

Annex A (informative) Rationale for requirements ...........................................................................................................................17

Annex B (informative) Risk management process for medical devices ...........................................................................26

Annex C (informative) Fundamental risk concepts ............................................................................................................................30

Bibliography .............................................................................................................................................................................................................................36

ISO 14971:2019(E)EVS-EN ISO 14971:2019

## Foreword

ISO (the International Organization for Standardization) is a worldwide federation of national standards bodies (ISO member bodies). The work of preparing International Standards is normally carried out through ISO technical committees. Each member body interested in a subject for which a technical committee has been established has the right to be represented on that committee. International organizations, governmental and non-governmental, in liaison with ISO, also take part in the work. ISO collaborates closely with the International Electrotechnical Commission (IEC) on all matters of electrotechnical standardization.

The procedures used to develop this document and those intended for its further maintenance are described in the ISO/IEC Directives, Part 1. In particular, the different approval criteria needed for the different types of ISO documents should be noted. This document was drafted in accordance with the editorial rules of the ISO/IEC Directives, Part 2 (see www .iso .org/ directives).

Attention is drawn to the possibility that some of the elements of this document may be the subject of patent rights. ISO shall not be held responsible for identifying any or all such patent rights. Details of any patent rights identified during the development of the document will be in the Introduction and/or on the ISO list of patent declarations received (see www .iso .org/ patents).

Any trade name used in this document is information given for the convenience of users and does not constitute an endorsement.

For an explanation of the voluntary nature of standards, the meaning of ISO specific terms and expressions related to conformity assessment, as well as information about ISO's adherence to the World Trade Organization (WTO) principles in the Technical Barriers to Trade (TBT) see  www .iso .org/ iso/ foreword .html.

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

This document was prepared by Technical Committee ISO/TC 210, Quality management and corresponding general aspects for medical devices, and IEC/SC 62A, Common aspects of electrical equipment used in medical practice.

This third edition cancels and replaces the second edition (ISO 14971:2007), which has been technically revised. The main changes compared to the previous edition are as follows:

— A clause on normative references has been included, in order to respect the requirements for fixed in Clause 15 of ISO/IEC Directives, Part 2:2018.

— The defined terms are updated and many are derived from ISO/IEC Guide 63:2019. Defined terms are printed in italic to assist the reader in identifying them in the body of the document.

— Definitions of benefit, reasonably foreseeable misuse and state of the art have been introduced.

— More attention is given to the benefits that are expected from the use of the medical device. The term benefit-risk analysis has been aligned with terminology used in some regulations.

— It is explained that the process described in ISO 14971 can be used for managing risks associated with medical devices, including those related to data and systems security.

— The method for the evaluation of the overall residual risk and the criteria for its acceptability are required to be defined in the risk management plan. The method can include gathering and reviewing data and literature for the medical device and for similar medical devices and similar other products on the market. The criteria for the acceptability of the overall residual risk can be different from the criteria for acceptability of individual risks.

— The requirements to disclose residual risks have been moved and merged into one requirement, after the overall residual risk has been evaluated and judged acceptable.

— The review before commercial distribution of the medical device concerns the execution of the risk management plan. The results of the review are documented as the risk management report.

iv 	 	© ISO 2019 – All rights reserved

EVS-EN ISO 14971:2019

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

— The requirements for production and post-production activities have been clarified and restructured. More detail is given on the information to be collected and the actions to be taken when the collected information has been reviewed and determined to be relevant to safety.

— Several informative annexes are moved to the guidance in ISO/TR 24971, which has been revised in parallel. More information and a rationale for the requirements in this third edition of ISO 14971 have been provided in Annex A. The correspondence between the clauses of the second edition and those of this third edition is given in Annex B.

Any feedback or questions on this document should be directed to the user’s national standards body. A complete listing of these bodies can be found at www .iso .org/ members .html.

ISO 14971:2019(E)EVS-EN ISO 14971:2019

## Introduction

The requirements contained in this document provide manufacturers with a framework within which experience, insight and judgment are applied systematically to manage the risks associated with the use of medical devices.

This document was developed specifically for manufacturers of medical devices on the basis of established principles of risk management that have evolved over many years. This document could be used as guidance in developing and maintaining a risk management process for other products that are not necessarily medical devices in some jurisdictions and for suppliers and other parties involved in the medical device life cycle.

This document deals with processes for managing risks associated with medical devices. Risks can be related to injury, not only to the patient, but also to the user and other persons. Risks can also be related to damage to property (for example objects, data, other equipment) or the environment.

Risk management is a complex subject because each stakeholder can place a different value on the acceptability of risks in relation to the anticipated benefits. The concepts of risk management are particularly important in relation to medical devices because of the variety of stakeholders including medical practitioners, the organizations providing health care, governments, industry, patients and members of the public.

It is generally accepted that the concept of risk has two key components:

— the probability of occurrence of harm; and

— the consequences of that harm, that is, how severe it might be.

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

All stakeholders need to understand that the use of a medical device involves an inherent degree of risk, even after the risks have been reduced to an acceptable level. It is well known that in the context of a clinical procedure some residual risks remain. The acceptability of a risk to a stakeholder is influenced by the key components listed above and by the stakeholder’s perception of the risk and the benefit. Each stakeholder’s perception can vary depending upon their cultural background, the socio-economic and educational background of the society concerned and the actual and perceived state of health of the patient. The way a risk is perceived also takes into account other factors, for example, whether exposure to the hazard or hazardous situation seems to be involuntary, avoidable, from a man-made source, due to negligence, arising from a poorly understood cause, or directed at a vulnerable group within society.

As one of the stakeholders, the manufacturer reduces risks and makes judgments relating to the safety of a medical device, including the acceptability of residual risks. The manufacturer takes into account the generally acknowledged state of the art, in order to determine the suitability of a medical device to be placed on the market for its intended use. This document specifies a process through which the manufacturer of a medical device can identify hazards associated with the medical device, estimate and evaluate the risks associated with these hazards, control these risks, and monitor the effectiveness of the controls throughout the life cycle of the medical device.

The decision to use a medical device in the context of a particular clinical procedure requires the residual risks to be balanced against the anticipated benefits of the procedure. Such decisions are beyond the scope of this document and take into account the intended use, the circumstances of use, the performance and risks associated with the medical device, as well as the risks and benefits associated with the clinical procedure. Some of these decisions can be made only by a qualified medical practitioner with knowledge of the state of health of an individual patient or the patient’s own opinion.

For any particular medical device, other standards or regulations could require the application of specific methods for managing risk. In those cases, it is necessary to also follow the requirements outlined in those documents.

vi 	 	© ISO 2019 – All rights reserved

EVS-EN ISO 14971:2019

The verbal forms used in this document conform to the usage described in Clause 7 of the ISO/ IEC Directives, Part 2:2018. For the purposes of this document, the auxiliary verb:

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

— “shall” means that compliance with a requirement or a test is mandatory for compliance with this document;

— “should” means that compliance with a requirement or a test is recommended but is not mandatory for compliance with this document;

— “may” is used to describe permission (e.g. a permissible way to achieve compliance with a requirement or test);

— “can” is used to express possibility and capability; and

— “must” is used to express an external constraint that is not a requirement of the document.

### INTERNATIONAL STANDARD 	EVS-EN ISO 14971:2019ISO 14971:2019(E)

Medical devices — Application of risk management to medical devices

### 1 	Scope

This document specifies terminology, principles and a process for risk management of medical devices, including software as a medical device and in vitro diagnostic medical devices. The process described in this document intends to assist manufacturers of medical devices to identify the hazards associated with the medical device, to estimate and evaluate the associated risks, to control these risks, and to monitor the effectiveness of the controls.

The requirements of this document are applicable to all phases of the life cycle of a medical device. The process described in this document applies to risks associated with a medical device, such as risks related to biocompatibility, data and systems security, electricity, moving parts, radiation, and usability.

The process described in this document can also be applied to products that are not necessarily medical devices in some jurisdictions and can also be used by others involved in the medical device life cycle.

This document does not apply to:

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

— decisions on the use of a medical device in the context of any particular clinical procedure; or — business risk management.

This document requires manufacturers to establish objective criteria for risk acceptability but does not specify acceptable risk levels.

Risk management can be an integral part of a quality management system. However, this document does not require the manufacturer to have a quality management system in place.

NOTE 	Guidance on the application of this document can be found in ISO/TR 24971[9].

2 	Normative references

There are no normative references in this document.

### 3	Terms	and	definitions

For the purposes of this document, the following terms and definitions apply.

ISO and IEC maintain terminological databases for use in standardization at the following addresses:

— ISO Online browsing platform: available at https:// www .iso .org/ obp

— IEC Electropedia: available at http:// www .electropedia .org/

#### 3.1 accompanying documentation

materials accompanying a medical device (3.10) and containing information for the user or those accountable for the installation, use, maintenance, decommissioning and disposal of the medical device (3.10), particularly regarding safe use

Note 1 to entry: The accompanying documentation can consist of the instructions for use, technical description, installation manual, quick reference guide, etc.

Note 2 to entry: Accompanying documentation is not necessarily a written or printed document but could involve auditory, visual, or tactile materials and multiple media types.

#### 3.2 benefit

positive impact or desirable outcome of the use of a medical device (3.10) on the health of an individual, or a positive impact on patient management or public health

Note 1 to entry: Benefits can include positive impact on clinical outcome, the patient’s quality of life, outcomes related to diagnosis, positive impact from diagnostic devices on clinical outcomes, or positive impact on public health.

3.3

#### harm

injury or damage to the health of people, or damage to property or the environment [SOURCE: ISO/IEC Guide 63:2019, 3.1]

3.4

#### hazard

potential source of harm (3.3)

[SOURCE: ISO/IEC Guide 63:2019, 3.2]

##### 3.5

hazardous situation circumstance in which people, property or the environment is/are exposed to one or more hazards (3.4)

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Note 1 to entry: See Annex C for an explanation of the relationship between hazard and hazardous situation.

[SOURCE: ISO/IEC Guide 63:2019, 3.3, modified — Note 1 to entry added.]

##### 3.6 intended use intended purpose

use for which a product, process (3.14) or service is intended according to the specifications, instructions and information provided by the manufacturer (3.9)

Note 1 to entry: The intended medical indication, patient population, part of the body or type of tissue interacted with, user profile, use environment, and operating principle are typical elements of the intended use. [SOURCE: ISO/IEC Guide 63:2019, 3.4]

3.7 in vitro diagnostic medical device IVD medical device device, whether used alone or in combination, intended by the manufacturer (3.9) for the in vitro examination of specimens derived from the human body solely or principally to provide information for diagnostic, monitoring or compatibility purposes and including reagents, calibrators, control materials, specimen receptacles, software, and related instruments or apparatus or other articles [SOURCE: ISO 18113-1:2009, 3.27, modified — NOTE deleted.]

3.8

#### life cycle

series of all phases in the life of a medical device (3.10), from the initial conception to final decommissioning and disposal

[SOURCE: ISO/IEC Guide 63:2019, 3.5]

##### 3.9 manufacturer

natural or legal person with responsibility for the design and/or manufacture of a medical device (3.10) with the intention of making the medical device (3.10) available for use, under his name, whether or not such a medical device (3.10) is designed and/or manufactured by that person himself or on his behalf by another person(s)

Note 1 to entry: The natural or legal person has ultimate legal responsibility for ensuring compliance with all applicable regulatory requirements for the medical device in the countries or jurisdictions where it is intended to be made available or sold, unless this responsibility is specifically imposed on another person by the Regulatory Authority (RA) within that jurisdiction.

Note 2 to entry: The manufacturer’s responsibilities are described in other GHTF guidance documents. These responsibilities include meeting both pre-market requirements and post-market requirements, such as adverse event reporting and notification of corrective actions.

Note 3 to entry: “Design and/or manufacture” may include specification development, production, fabrication, assembly, processing, packaging, repackaging, labelling, relabelling, sterilization, installation, or remanufacturing of a medical device; or putting a collection of devices, and possibly other products, together for a medical purpose.

Note 4 to entry: Any person who assembles or adapts a medical device that has already been supplied by another person for an individual patient, in accordance with the instructions for use, is not the manufacturer, provided the assembly or adaptation does not change the intended use of the medical device.

Note 5 to entry: Any person who changes the intended use of, or modifies, a medical device without acting on behalf of the original manufacturer and who makes it available for use under his own name, should be considered the manufacturer of the modified medical device.

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Note 6 to entry: An authorised representative, distributor or importer who only adds its own address and contact details to the medical device or the packaging, without covering or changing the existing labelling, is not considered a manufacturer.

Note 7 to entry: To the extent that an accessory is subject to the regulatory requirements of a medical device, the person responsible for the design and/or manufacture of that accessory is considered to be a manufacturer. [SOURCE: ISO/IEC Guide 63:2019, 3.6]

3.10 medical device instrument, apparatus, implement, machine, appliance, implant, reagent for in vitro use, software, material or other similar or related article, intended by the manufacturer (3.9) to be used, alone or in combination, for human beings, for one or more of the specific medical purpose(s) of — diagnosis, prevention, monitoring, treatment or alleviation of disease,

— diagnosis, monitoring, treatment, alleviation of or compensation for an injury,

— investigation, replacement, modification, or support of the anatomy or of a physiological process,

— supporting or sustaining life,

— control of conception,

— disinfection of medical devices (3.10),

— providing information by means of in vitro examination of specimens derived from the human body,

and which does not achieve its primary intended action by pharmacological, immunological or metabolic means, in or on the human body, but which may be assisted in its function by such means

Note 1 to entry: Products which can be considered to be medical devices in some jurisdictions but not in others include:

— 	disinfection substances;

— 	aids for persons with disabilities;

— 	devices incorporating animal and/or human tissues;

— 	devices for in vitro fertilization or assisted reproduction technologies.

[SOURCE: ISO/IEC Guide 63:2019, 3.7]

3.11 objective evidence data supporting the existence or verity of something

Note 1 to entry: Objective evidence can be obtained through observation, measurement, test or by other means.

[SOURCE: ISO 9000:2015, 3.8.3, modified — Note 2 to entry deleted.]

##### 3.12 post-production

part of the life cycle (3.8) of the medical device (3.10) after the design has been completed and the medical device (3.10) has been manufactured

EXAMPLE Transportation, storage, installation, product use, maintenance, repair, product changes, decommissioning and disposal.

3.13

#### procedure

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

specified way to carry out an activity or a process (3.14)

Note 1 to entry: Procedures can be documented or not.

[SOURCE: ISO 9000:2015, 3.4.5]

3.14 process set of interrelated or interacting activities that use inputs to deliver an intended result

Note 1 to entry: Whether the “intended result” of a process is called output, product or service depends on the context of the reference.

Note 2 to entry: Inputs to a process are generally the outputs of other processes and outputs of a process are generally the inputs to other processes.

Note 3 to entry: Two or more interrelated and interacting processes in series can also be referred to as a process.

[SOURCE: ISO 9000:2015, 3.4.1, modified — Notes to entry 4, 5 and 6 are deleted.]

##### 3.15 reasonably foreseeable misuse

use of a product or system in a way not intended by the manufacturer (3.9), but which can result from readily predictable human behaviour

Note 1 to entry: Readily predictable human behaviour includes the behaviour of all types of users, e.g. lay and professional users.

Note 2 to entry: Reasonably foreseeable misuse can be intentional or unintentional.

[SOURCE: ISO/IEC Guide 63:2019, 3.8]

3.16 record document stating results achieved or providing evidence of activities performed

Note 1 to entry: Records can be used, for example, to formalize traceability and to provide evidence of verification, preventive action and corrective action.

Note 2 to entry: Generally records need not be under revision control.

[SOURCE: ISO 9000:2015, 3.8.10]

##### 3.17 residual risk

risk remaining after risk control (3.21) measures have been implemented [SOURCE: ISO/IEC Guide 63:2019, 3.9]

##### 3.18 risk

combination of the probability of occurrence of harm (3.3) and the severity (3.27) of that harm (3.3)

[SOURCE: ISO/IEC Guide 63:2019, 3.10, modified — Note 1 to entry deleted.]

##### 3.19 risk analysis

systematic use of available information to identify hazards (3.4) and to estimate the risk (3.18) [SOURCE: ISO/IEC Guide 63:2019, 3.11]

##### 3.20 risk assessment

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

overall process (3.14) comprising a risk analysis (3.19) and a risk evaluation (3.20) [SOURCE: ISO/IEC Guide 51:2014, 3.11]

##### 3.21 risk control

process (3.14) in which decisions are made and measures implemented by which risks (3.18) are reduced to, or maintained within, specified levels

[SOURCE: ISO/IEC Guide 63:2019, 3.12]

##### 3.22 risk estimation

process (3.14) used to assign values to the probability of occurrence of harm (3.3) and the severity (3.27) of that harm

[SOURCE: ISO/IEC Guide 63:2019, 3.13]

##### 3.23 risk evaluation

process (3.14) of comparing the estimated risk (3.18) against given risk (3.18) criteria to determine the acceptability of the risk (3.18)

[SOURCE: ISO/IEC Guide 63:2019, 3.14]

##### 3.24 risk management

systematic application of management policies, procedures (3.13) and practices to the tasks of analysing, evaluating, controlling and monitoring risk (3.18)

[SOURCE: ISO/IEC Guide 63:2019, 3.15]

##### 3.25 risk management file

set of records (3.16) and other documents that are produced by risk management (3.24)

3.26 safety freedom from unacceptable risk (3.18)

[SOURCE: ISO/IEC Guide 63:2019, 3.10]

##### 3.27 severity

measure of the possible consequences of a hazard (3.4)

[SOURCE: ISO/IEC Guide 63:2019, 3.17]

3.28

#### state of the art

developed stage of technical capability at a given time as regards products, processes (3.14) and services, based on the relevant consolidated findings of science, technology and experience

Note 1 to entry: The state of the art embodies what is currently and generally accepted as good practice in technology and medicine. The state of the art does not necessarily imply the most technologically advanced solution. The state of the art described here is sometimes referred to as the “generally acknowledged state of the art”. [SOURCE: ISO/IEC Guide 63:2019, 3.18]

##### 3.29 top management

person or group of people who directs and controls a manufacturer (3.9) at the highest level

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

[SOURCE: ISO 9000:2015, 3.1.1, modified — “An organization” replaced by “a manufacturer”, Notes to entry deleted.]

##### 3.30 use error

user action or lack of user action while using the medical device (3.10) that leads to a different result than that intended by the manufacturer (3.9) or expected by the user

Note 1 to entry: Use error includes the inability of the user to complete a task.

Note 2 to entry: Use errors can result from a mismatch between the characteristics of the user, user interface, task, or use environment.

Note 3 to entry: Users might be aware or unaware that a use error has occurred.

Note 4 to entry: An unexpected physiological response of the patient is not by itself considered use error.

Note 5 to entry: A malfunction of a medical device that causes an unexpected result is not considered a use error.

[SOURCE: IEC 62366-1:2015, 3.21, modified — Note 6 to entry deleted.]

##### 3.31 verification

confirmation, through the provision of objective evidence (3.11), that specified requirements have been fulfilled

Note 1 to entry: The objective evidence needed for a verification can be the result of an inspection or of other forms of determination such as performing alternative calculations or reviewing documents.

Note 2 to entry: The activities carried out for verification are sometimes called a qualification process.

Note 3 to entry: The word “verified” is used to designate the corresponding status.

[SOURCE: ISO/IEC Guide 63:2019, 3.19]

### 4 	General requirements for risk management system

#### 4.1 	Risk management process

The manufacturer shall establish, implement, document and maintain an ongoing process for:

- identifying hazards and hazardous situations associated with a medical device;

- estimating and evaluating the associated risks;

- controlling these risks, and

- monitoring the effectiveness of the risk control measures.

This process shall apply throughout the life cycle of the medical device.

This process shall include the following elements:

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

— risk analysis;

— risk evaluation;

— risk control; and

— production and post-production activities.

Where a documented product realization process exists, it shall incorporate the appropriate parts of the risk management process.

NOTE 1 	Product realization processes are described in, for example, Clause 7 of ISO 13485:2016[5].

NOTE 2 A documented process within a quality management system can be used to address safety in a systematic manner, in particular to enable the early identification of hazards and hazardous situations in complex medical devices.

NOTE 3 A schematic representation of the risk management process is shown in Figure 1. Depending on the specific life cycle phase, individual elements of risk management can have varying emphasis. Also, risk management activities can be performed iteratively or in multiple steps as appropriate to the medical device. Annex B contains a more detailed overview of the steps in the risk management process.

Compliance is checked by inspection of the appropriate documents.

Figure 1 — A schematic representation of the risk management process

#### 4.2 	Management responsibilities

Top management shall provide evidence of its commitment to the risk management process by ensuring:

— the provision of adequate resources; and

— the assignment of competent personnel (see 4.3) for risk management.

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Top management shall define and document a policy for establishing criteria for risk acceptability. The policy shall provide a framework that ensures that criteria are based upon applicable national or regional regulations and relevant International Standards, and take into account available information such as the generally acknowledged state of the art and known stakeholder concerns.

NOTE 1 The manufacturer’s policy for establishing criteria for risk acceptability can define the approaches to risk control: reducing risk as low as reasonably practicable, reducing risk as low as reasonably achievable, or reducing risk as far as possible without adversely affecting the benefit-risk ratio. See ISO/TR 24971[9] for guidance on defining such policy.

Top management shall review the suitability of the risk management process at planned intervals to ensure continuing effectiveness of the risk management process and shall document any decisions and actions taken. If the manufacturer has a quality management system in place, this review may be part of the quality management system review.

NOTE 2 The results of reviewing production and post-production information can be an input to the review of the suitability of the risk management process.

NOTE 3 The documents described in this subclause can be incorporated within the documents produced by the manufacturer’s quality management system and these documents can be referenced in the risk management file.

Compliance is checked by inspection of the appropriate documents.

#### 4.3 	Competence of personnel

Persons performing risk management tasks shall be competent on the basis of education, training, skills and experience appropriate to the tasks assigned to them. Where appropriate, these persons shall have knowledge of and experience with the particular medical device (or similar medical devices) and its use, the technologies involved or the risk management techniques employed. Appropriate records shall be maintained.

NOTE Risk management tasks can be performed by representatives of several functions, each contributing their specialist knowledge.

Compliance is checked by inspection of the appropriate records.

#### 4.4 	Risk management plan

Risk management activities shall be planned. For the particular medical device being considered, the manufacturer shall establish and document a risk management plan in accordance with the risk management process. The risk management plan shall be part of the risk management file.

This plan shall include at least the following:

- the scope of the planned risk management activities, identifying and describing the medical device and the life cycle phases for which each element of the plan is applicable;

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

assignment of responsibilities and authorities;

- requirements for review of risk management activities;

- criteria for risk acceptability, based on the manufacturer’s policy for determining acceptable risk, including criteria for accepting risks when the probability of occurrence of harm cannot be estimated;

NOTE 1 The criteria for risk acceptability are essential for the ultimate effectiveness of the risk management process. For each risk management plan the manufacturer needs to establish risk acceptability criteria that are appropriate for the particular medical device.

- a method to evaluate the overall residual risk, and criteria for acceptability of the overall residual risk based on the manufacturer’s policy for determining acceptable risk;

NOTE 2 The method to evaluate the overall residual risk can include gathering and reviewing data and literature for the medical device being considered and similar medical devices on the market and can involve judgment by a cross-functional team of experts with application knowledge and clinical expertise.

- activities for verification of the implementation and effectiveness of risk control measures; and

- activities related to collection and review of relevant production and post-production information.

NOTE 3 See ISO/TR 24971[9] for guidance on developing a risk management plan and on establishing criteria for risk acceptability.

NOTE 4 Not all parts of the plan need to be created at the same time. The plan or parts of it can be developed over time.

If the plan changes during the life cycle of the medical device, a record of the changes shall be maintained in the risk management file.

Compliance is checked by inspection of the risk management file.

#### 4.5 	Risk management file

For the particular medical device being considered, the manufacturer shall establish and maintain a risk management file. In addition to the requirements of other clauses of this document, the risk management file shall provide traceability for each identified hazard to:

— the risk analysis;

— the risk evaluation;

— the implementation and verification of the risk control measures; and — the results of the evaluation of the residual risks.

NOTE 1 The records and other documents that make up the risk management file can form part of other documents and files required, for example, by a manufacturer’s quality management system. The risk management file need not physically contain all the records and other documents. However, it needs to contain at least references or pointers to all required documentation, so that the manufacturer can assemble the information referenced in the risk management file in a timely manner.

NOTE 2 	The risk management file can be in any form or type of medium.

NOTE 3 See ISO/TR 24971[9] for guidance on establishing a risk management file for components and devices that were designed without using ISO 14971.

### 5 	Risk analysis

#### 5.1 	Risk analysis process

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

The manufacturer shall perform risk analysis for the particular medical device as described in 5.2 to 5.5. The implementation of the planned risk analysis activities and the results of the risk analysis shall be recorded in the risk management file.

NOTE 1 If a risk analysis or other relevant information is available for a similar medical device, that analysis or information can be used as a starting point for the new risk analysis. The degree of relevance depends on the differences between the medical devices and whether these introduce new hazards or significant differences in outputs, characteristics, performance or results. The extent of use of an existing risk analysis is based on a systematic evaluation of the effects that the differences can have on the occurrence of hazardous situations.

NOTE 2 See ISO/TR 24971[9] for guidance on selected risk analysis techniques and on risk analysis techniques for in vitro diagnostic medical devices.

In addition to the records required in 5.2 to 5.5, the documentation of the conduct and results of the risk analysis shall include at least the following:

- identification and description of the medical device that was analysed;

- identification of the person(s) and organization who carried out the risk analysis; and c) 	scope and date of the risk analysis.

NOTE 3 The scope of the risk analysis can be very broad (as for the development of a new medical device with which a manufacturer has little or no experience) or the scope can be limited (as for analysing the impact of a change to an existing medical device for which much information already exists in the manufacturer’s files).

Compliance is checked by inspection of the risk management file.

#### 5.2 	Intended use and reasonably foreseeable misuse

The manufacturer shall document the intended use of the particular medical device being considered.

The intended use should take into account information such as the intended medical indication, patient population, part of the body or type of tissue interacted with, user profile, use environment, and operating principle.

The manufacturer shall also document reasonably foreseeable misuse.

This documentation shall be maintained in the risk management file.

NOTE 1 	The use specification (see 3.23 of IEC 62366-1:2015[13]) can be an input to determining the intended use.

NOTE 2 See ISO/TR 24971[9] for factors to consider in determining the intended use and for an explanation of reasonably foreseeable misuse.

Compliance is checked by inspection of the risk management file.

#### 5.3	Identification	of	characteristics	related	to	safety

For the particular medical device being considered, the manufacturer shall identify and document those qualitative and quantitative characteristics that could affect the safety of the medical device. Where appropriate, the manufacturer shall define limits of those characteristics. This documentation shall be maintained in the risk management file.

NOTE 1 See ISO/TR 24971[9] for a list of questions that can serve as a guide in identifying medical device characteristics that could have an impact on safety.

NOTE 2 Characteristics related to loss or degradation of the clinical performance of a medical device that can result in unacceptable risk, are sometimes referred to as essential performance (see for example IEC 60601-1[12]).

Compliance is checked by inspection of the risk management file.

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

5.4	Identification	of	hazards and hazardous situations

The manufacturer shall identify and document known and foreseeable hazards associated with the medical device based on the intended use, reasonably foreseeable misuse and the characteristics related to safety in both normal and fault conditions.

For each identified hazard, the manufacturer shall consider the reasonably foreseeable sequences or combinations of events that can result in a hazardous situation, and shall identify and document the resulting hazardous situation(s).

NOTE 1 A sequence of events can be initiated in all phases of the life cycle, e.g. during transport, storage, installation, maintenance, routine inspection, decommissioning and disposal.

NOTE 2 An explanation of the relationship between hazard, hazardous situation and harm including examples is given in Annex C.

NOTE 3 Risk analysis includes the examination of different sequences or combinations of events related to a single hazard that can lead to different hazardous situations. Each hazardous situation can lead to different types of harm.

NOTE 4 When identifying hazardous situations not previously recognised, systematic techniques for risk analysis that cover the specific situation can be used. Guidance on some available techniques is provided in ISO/TR 24971[9].

The documentation shall be maintained in the risk management file.

Compliance is checked by inspection of the risk management file.

#### 5.5 	Risk estimation

For each identified hazardous situation, the manufacturer shall estimate the associated risk(s) using available information or data. For hazardous situations for which the probability of the occurrence of harm cannot be estimated, the possible consequences shall be listed for use in risk evaluation and risk control. The results of these activities shall be recorded in the risk management file.

The system used for qualitative or quantitative categorization of probability of occurrence of harm and severity of harm shall be recorded in the risk management file.

NOTE 1 Risk estimation incorporates an analysis of the probability of occurrence of harm and the severity of the harm. Depending on the area of application, only certain elements of the risk estimation process might need to be considered in detail. For example, when the harm is minimal, an initial hazard and consequence analysis could be sufficient, or when insufficient information or data are available, a conservative estimate of the probability of occurrence can give some indication of the risk. See also ISO/TR 24971[9].

NOTE 2 Risk estimation can be qualitative or quantitative. Methods of risk estimation, including those resulting from systematic faults, are described in ISO/TR 24971[9], which also gives information useful for estimating risks for in vitro diagnostic medical devices.

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

NOTE 3 	Information or data for estimating risks can be obtained, for example, from:

| — | published standards; |
| --- | --- |
| — | scientific or technical investigations; |
| — | field data from similar medical devices already in use, including publicly available reports of incidents; |
| — | usability tests employing typical users; |
| — | clinical evidence; |
| — | results of relevant investigations or simulations; |
| — | expert opinion; or |
| — | external quality assessment schemes for in vitro diagnostic medical devices. |

Compliance is checked by inspection of the risk management file.

### 6 	Risk evaluation

For each identified hazardous situation, the manufacturer shall evaluate the estimated risks and determine if the risk is acceptable or not, using the criteria for risk acceptability defined in the risk management plan.

If the risk is acceptable, it is not required to apply the requirements given in 7.1 to 7.5 to this hazardous situation (i.e., proceed to 7.6) and the estimated risk shall be treated as residual risk.

If the risk is not acceptable, then the manufacturer shall perform risk control activities as described in 7.1 to 7.6.

The results of this risk evaluation shall be recorded in the risk management file.

Compliance is checked by inspection of the risk management file.

### 7 	Risk control

#### 7.1 	Risk control option analysis

The manufacturer shall determine risk control measures that are appropriate for reducing the risks to an acceptable level.

The manufacturer shall use one or more of the following risk control options in the priority order listed: a) 	inherently safe design and manufacture;

- protective measures in the medical device itself or in the manufacturing process;

- information for safety and, where appropriate, training to users.

NOTE 1 	The rationale for the priority order in selecting the risk control options is given in A.2.7.1.

NOTE 2 Risk control measures can reduce the severity of the harm or reduce the probability of occurrence of the harm, or both.

NOTE 3 	See ISO/TR 24971[9] for guidance on providing information for safety.

Relevant standards should be applied as part of the risk control option analysis.

NOTE 4 Many standards address inherent safety, protective measures, and information for safety for medical devices. In addition, some medical device standards have integrated elements of the risk management process (e.g. electromagnetic compatibility, usability, biological evaluation). See ISO/TR 24971[9] for information on the role of International Standards in risk management.

The risk control measures selected shall be recorded in the risk management file.

If, during risk control option analysis, the manufacturer determines that risk reduction is not practicable, the manufacturer shall conduct a benefit-risk analysis of the residual risk (proceed to 7.4).

Compliance is checked by inspection of the risk management file.

#### 7.2 	Implementation of risk control measures

The manufacturer shall implement the risk control measures selected in 7.1.

Implementation of each risk control measure shall be verified. This verification shall be recorded in the risk management file.

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

NOTE 1 Verification of implementation can be performed as part of design and development verification or process qualification within a quality management system.

The effectiveness of the risk control measures shall be verified. The results of this verification shall be recorded in the risk management file.

NOTE 2 Verification of effectiveness can be performed as part of design and development validation within a quality management system and can include testing with users. See A.2.7.2.

NOTE 3 Verification of effectiveness can also be performed as part of design and development verification or process qualification, if the relationship between the effectiveness in risk reduction and the result of design and development verification or process qualification is known.

EXAMPLE 1 Design verification of a certain performance characteristic, such as dose accuracy of a drug injector, can serve as verification of effectiveness of risk control measures ensuring safe drug dosing.

EXAMPLE 2 Process qualification can serve as verification of effectiveness of risk control measures related to risk caused by variations in production output.

NOTE 4 See ISO 13485[5] for more information on design and development verification and validation. See also ISO/TR 24971[9] for more guidance.

Compliance is checked by inspection of the risk management file.

#### 7.3 	Residual risk evaluation

After the risk control measures are implemented, the manufacturer shall evaluate the residual risk using the criteria for risk acceptability defined in the risk management plan. The results of this evaluation shall be recorded in the risk management file.

If a residual risk is not judged acceptable using these criteria, further risk control measures shall be considered (go back to 7.1).

Compliance is checked by inspection of the risk management file.

#### 7.4 	Benefit-risk analysis

If a residual risk is not judged acceptable using the criteria established in the risk management plan and further risk control is not practicable, the manufacturer may gather and review data and literature to determine if the benefits of the intended use outweigh this residual risk.

If this evidence does not support the conclusion that the benefits outweigh this residual risk, then the manufacturer may consider modifying the medical device or its intended use (go back to 5.2). Otherwise, this risk remains unacceptable.

If the benefits outweigh the residual risk, then proceed to 7.5.

The results of the benefit-risk analysis shall be recorded in the risk management file.

NOTE 	See ISO/TR 24971[9] for guidance on performing a benefit-risk analysis.

Compliance is checked by inspection of the risk management file.

#### 7.5 	Risks arising from risk control measures

The manufacturer shall review the effects of the risk control measures with regard to whether:

— new hazards or hazardous situations are introduced; or

— the estimated risks for previously identified hazardous situations are affected by the introduction of the risk control measures.

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Any new or increased risks shall be managed in accordance with 5.5 to 7.4.

The results of this review shall be recorded in the risk management file.

Compliance is checked by inspection of the risk management file.

#### 7.6 	Completeness of risk control

The manufacturer shall review the risk control activities to ensure that the risks from all identified hazardous situations have been considered and all risk control activities are completed.

The results of this review shall be recorded in the risk management file.

Compliance is checked by inspection of the risk management file.

### 8 	Evaluation of overall residual risk

After all risk control measures have been implemented and verified, the manufacturer shall evaluate the overall residual risk posed by the medical device, taking into account the contributions of all residual risks, in relation to the benefits of the intended use, using the method and the criteria for acceptability of the overall residual risk defined in the risk management plan [see 4.4 e)].

If the overall residual risk is judged acceptable, the manufacturer shall inform users of significant residual risks and shall include the necessary information in the accompanying documentation in order to disclose those residual risks.

NOTE 1 	The rationale for the disclosure of significant residual risks is given in A.2.8.

NOTE 2 See ISO/TR 24971[9] for guidance on the evaluation of overall residual risk and the disclosure of residual risks.

If the overall residual risk is not judged acceptable in relation to the benefits of the intended use, the manufacturer may consider implementing additional risk control measures (go back to 7.1) or modifying the medical device or its intended use (go back to 5.2). Otherwise, the overall residual risk remains unacceptable.

The results of the evaluation of the overall residual risk shall be recorded in the risk management file.

Compliance is checked by inspection of the risk management file and the accompanying documentation.

### 9 	Risk management review

Prior to release for commercial distribution of the medical device, the manufacturer shall review the execution of the risk management plan. This review shall at least ensure that:

— the risk management plan has been appropriately implemented;

— the overall residual risk is acceptable; and

— appropriate methods are in place to collect and review information in the production and postproduction phases.

The results of this review shall be recorded and maintained as the risk management report and shall be included in the risk management file.

The responsibility for review shall be assigned in the risk management plan to persons having the appropriate authority [see 4.4 b)].

Compliance is checked by inspection of the risk management file.

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

10 Production and post-production activities

#### 10.1 General

The manufacturer shall establish, document and maintain a system to actively collect and review information relevant to the medical device in the production and post-production phases. When establishing this system, the manufacturer shall consider appropriate methods for the collection and processing of information.

NOTE 1 	See also 7.3.3, 8.2.1, 8.4 and 8.5 of ISO 13485:2016[5].

NOTE 2 	See ISO/TR 24971[9] for guidance on production and post-production activities.

Compliance is checked by inspection of the appropriate documents.

#### 10.2 Information collection

The manufacturer shall collect, where applicable:

- information generated during production and monitoring of the production process;

- information generated by the user;

- information generated by those accountable for the installation, use and maintenance of the medical device;

- information generated by the supply chain;

- publicly available information; and

- information related to the generally acknowledged state of the art.

NOTE Information related to the generally acknowledged state of the art can include new or revised standards, published validated data specific to the application of the medical device under consideration, the availability of alternative medical devices and/or therapies, and other information (see also ISO/TR 24971[9]).

The manufacturer shall also consider the need to actively collect and review publicly available information about similar medical devices and similar other products on the market.

Compliance is checked by inspection of the appropriate documents.

#### 10.3 Information review

The manufacturer shall review the information collected for possible relevance to safety, especially whether:

— previously unrecognised hazards or hazardous situations are present;

— an estimated risk arising from a hazardous situation is no longer acceptable;

— the overall residual risk is no longer acceptable in relation to the benefits of the intended use; or — the generally acknowledged state of the art has changed.

The results of the review shall be recorded in the risk management file.

Compliance is checked by inspection of the risk management file.

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

10.4 Actions

If the collected information is determined to be relevant to safety, the following actions apply.

- Concerning the particular medical device,

— the manufacturer shall review the risk management file and determine if reassessment of risks and/or assessment of new risks is necessary;

— if a residual risk is no longer acceptable, the impact on previously implemented risk control measures shall be evaluated and should be considered as an input for modification of the medical device;

— the manufacturer should consider the need for actions regarding medical devices on the market; and

— any decisions and actions shall be recorded in the risk management file.

- Concerning the risk management process,

— the manufacturer shall evaluate the impact on previously implemented risk management activities; and

— the results of this evaluation shall be considered as an input for the review of the suitability of the risk management process by top management (see 4.2).

NOTE Some aspects of post-production monitoring are the subject of some national regulations. In such cases, additional measures might be required (e.g. prospective post-production evaluations).

Compliance is checked by inspection of the risk management file and other appropriate documents.

Annex A

(informative)

## Rationale for requirements

### A.1 General

The ISO/TC 210 — IEC/SC 62A Joint Working Group 1 (JWG 1), Application of risk management to medical devices, developed this rationale to document its reasoning for establishing the various requirements contained in this document. Those who make future revisions can use this annex, along with experience gained in the use of this document, to make this document more useful to manufacturers, regulatory bodies and health care providers.

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

ISO Technical Committee 210 and IEC Subcommittee 62A decided to combine their efforts on risk management and to form JWG 1 with the task to develop a standard for the application of risk management to medical devices. When discussions on an International Standard for risk management began, crucial features of risk management needed to be addressed, such as the process of risk evaluation as well as the balancing of risks and benefits for medical devices. Manufacturers, regulatory bodies, and health care providers had recognised that “absolute safety” in medical devices was not achievable. In addition, the risks that derive from the increasing diversity of medical devices and their applications cannot be completely addressed through product safety standards. The recognition of these facts and the consequent need to manage risks from medical devices throughout their life cycle led to the decision to develop ISO 14971 as a tool to actively improve the safety of medical devices. The first edition of this standard was published in 2000.

The second edition of ISO 14971 was developed and published in 2007 to address the need for additional guidance on its application and on the relationship between hazards and hazardous situations. Minor changes were made to the normative section, such as the addition of the requirement to plan for post-production monitoring and the removal of the requirement for traceability from the risk management report.

The systematic review in 2010 revealed the need for further guidance on a few specific topics. It was decided to develop the technical report ISO/TR 24971[9], because even a small update of the guidance would necessitate a revision of the standard. The first edition of this report was published in 2013.

This third edition was developed to clarify the normative requirements and to describe them in more detail, in particular the clauses on the evaluation of overall residual risk, on the risk management review and report and on production and post-production information. The clarifications were deemed necessary in view of requests for explanation in the systematic review of ISO 14971 in 2016 and in view of stricter requirements from regulatory bodies. More emphasis was put on the benefits that are anticipated from the use of the medical device and the balance between the (overall) residual risks and those benefits. It was explained that the process described in ISO 14971 can be applied to all types of hazards and risks associated with a medical device, for example biocompatibility, data and systems security, electricity, moving parts, radiation or usability. Several informative annexes were moved from this document to the guidance in ISO/TR 24971, which was revised in parallel. This allows for more frequent updates of the guidance independent of revising the standard.

### A.2 Rationale for requirements in particular clauses and subclauses

#### A.2.1 Scope

As explained in the introduction to this document, a risk management standard applying to the life cycle of medical devices is required. Software as a medical device and in vitro diagnostic medical devices are specifically mentioned in the scope to avoid any misunderstanding that, due to different regulations, these devices might be excluded from this document.

Risks can be present throughout the life cycle of the medical device, and risks that become apparent at one point in the life cycle can be managed by action taken at a completely different point in the life cycle. For this reason, the standard needs to be a complete life cycle standard. This means that the standard instructs manufacturers to apply risk management principles to a medical device from its initial conception until its ultimate decommissioning and disposal.

The process described in ISO 14971 can be applied to hazards and risks associated with the medical device. Risks related to data and systems security are specifically mentioned in the scope, to avoid any misunderstanding that a separate process would be needed to manage security risks related to medical devices. This does not preclude the possibility of developing specific standards, in which specific methods and requirements are provided for the assessment and control of security risks. Such standards can be used in conjunction with ISO 14971, in a similar way as IEC 62366-1[13] for usability, ISO 10993-1[4] for biological evaluation, or IEC 60601-1[12] for electrical and mechanical risks.

The scope of this document does not include clinical decision making, i.e., decisions on the use of a medical device in the context of a particular clinical procedure. Such decisions require the residual risks to be balanced against the anticipated benefits of the procedure or the risks and anticipated benefits of alternative procedures. Such decisions take into account the intended use, performance and risks associated with the medical device as well as the risks and benefits associated with the clinical procedure or the circumstances of use. Some of these decisions can be made only by a qualified health care professional with knowledge of the state of health of an individual patient and the patient’s own opinion.

The scope of this document also does not include business decision making. Other standards such as ISO 31000[10] exist for organisational risk management and related topics.

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Although there has been significant debate over what constitutes an acceptable level of risk, this document does not specify acceptability levels. Specifying a universal level for acceptable risk could be inappropriate. This decision is based upon the belief that:

— the wide variety of medical devices and situations covered by this document would make a universal level for acceptable risk meaningless;

— local laws, customs, values and perception of risk are more appropriate for defining risk acceptability for a particular culture or region of the world.

Because not all countries require a quality management system for medical device manufacturers, a quality management system is not a requirement of this document. However, a quality management system is extremely helpful in managing risks properly. Because of this and because most medical device manufacturers do employ a quality management system, this document is constructed so that it can easily be incorporated into the quality management system that they use.

#### A.2.2 Normative references

No other standards are required in order to establish and maintain a risk management process in accordance with ISO 14971. Clause 15 of ISO/IEC Directives, Part 2:2018, requires standards to include this statement.

#### A.2.3	Terms	and	definitions

Most of the definitions used in this document are taken from ISO 9000:2015[3] and ISO/IEC Guide 63:2019[2] which in turn adopted and adapted many of the definitions in ISO/IEC Guide 51:2014[1] and the definitions developed by the Global Harmonization Task Force (GHTF). Some of these definitions have a slightly different meaning in ISO/IEC Guide 63:2019[2] and ISO 14971 than in other standards.

For example, JWG 1 intended the definition of harm (3.3) to have a broad range and to include unreasonable psychological stress or unwanted pregnancy as part of “damage to the health of people”. Such stress can occur after a false positive diagnosis of a disease. “Damage to property and the environment” is undesirable and the associated risks need to be considered as well, for example those related to hazardous waste materials created by the use or disposal of the medical device. The word “physical” is removed from the definition of harm in ISO/IEC Guide 51:2014[1] and thus also in ISO/IEC Guide 63:2019 [2] and this document, because injury by itself already includes physical damage. Breaches of data and systems security can lead to harm, e.g. through loss of data, uncontrolled access to data, corruption or loss of diagnostic information, or corruption of software leading to malfunction of the medical device.

The definition of the term intended use (3.6) combines the definition of intended use as used in the United States and intended purpose which is the term in the European Union. These terms have essentially the same definition. It was intended that, when determining the intended use of a medical device, the manufacturer takes into account the intended medical indication, patient population, part of the body or tissue interacted with, user profile, use environment, and operating principle. The definition of life cycle (3.8) was necessary to make it clear that the term as used in this document covers all aspects of the existence of a medical device. The definition for risk management (3.24) emphasises the use of a systematic approach and the need for management oversight. The definition of top management (3.29) uses the definition from ISO 9000:2015[3]. It applies to the person or group at the highest level in the manufacturer’s organization.

Three other terms in ISO 14971 are not based on definitions in ISO/IEC Guide 63:2019[2] or in other standards. They are benefit (3.2), post-production (3.12) and risk management file (3.25). The term benefit is defined because of the increased emphasis by regulatory bodies on balancing the (residual) risks against the benefits of the medical device. For the same reason the phrase “benefit-risk analysis” is used. A definition of post-production was added to emphasise that the entire life cycle of the medical device is important for risk management. The concept of a risk management file is now well understood.

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

A.2.4 General requirements for risk management system

##### A.2.4.1 	Risk management process

The risk management system consists of the elements in 4.1 through 4.5.

The manufacturer needs to establish a risk management process as part of the design and development of a medical device. This is required so that the manufacturer can systematically ensure that the required elements are in the process. Risk analysis, risk evaluation and risk control are commonly recognised as essential parts of risk management. In addition to these elements, this document emphasises that the risk management process does not end with the design and production (including, as relevant, sterilization, packaging, and labelling) of a medical device, but continues on into the post-production phase. Therefore, the collection and review of production and post-production information was identified as a required part of the risk management process. Furthermore, it was felt that when a manufacturer employs a quality management system, the risk management process should be fully integrated into that quality management system.

Although risk management activities are highly individual to the medical device being considered, there are basic elements that need to be included in the risk management process. This need is addressed in 4.1. This subclause also recognises that there can be some differences in regulatory approaches to applying risk management to medical devices.

Subclauses 4.2 and 4.3 closely follow the risk-related requirements of quality management system standards. In some countries a quality management system is always required to market a medical device (unless the medical device is specifically exempted). In other countries manufacturers can choose whether to apply a quality management system. However, the requirements of 4.2 and 4.3 are always needed for an effective risk management process, whether or not the manufacturer operates all the other elements of a quality management system.

##### A.2.4.2 	Management responsibilities

The commitment of top management is critical for an effective risk management process. These individuals are responsible for overall guidance of the risk management process and this subclause is intended to emphasise their role. In particular:

— in the absence of adequate resources, risk management activities would be less effective, even if complying, to the letter, with the other requirements of this document;

— risk management is a specialized discipline and requires the involvement of competent individuals trained in risk management techniques (see A.2.4.3);

— because this document does not define acceptable risk levels, top management is required to establish a policy on how acceptable risks will be determined;

— risk management is an evolving process and periodic review of the risk management activities is needed to ascertain whether they are being carried out correctly, to rectify any weaknesses, to implement improvements, and to adapt to changes.

##### A.2.4.3 	Competence of personnel

It is most important to get competent people with the knowledge and experience necessary to perform risk management tasks. The risk management process requires people with knowledge and experience in areas such as:

— how the medical device is constructed;

— how the medical device works;

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

— how the medical device is produced; — how the medical device is actually used;

— how to apply the risk management process.

In general, this usually requires several representatives from various functions or disciplines, each contributing their specialist knowledge. The balance and relation between those representatives should be considered.

Records are required to provide objective evidence of competence. In order to avoid duplication and because of confidentiality and data protection considerations, this document does not require these records to be kept in the risk management file.

###### A.2.4.4 	Risk management plan

A risk management plan is required because:

— an organized approach is essential for good risk management;

— the plan provides the roadmap for risk management;

— the plan encourages objectivity and helps prevent essential elements being forgotten.

The elements a) to g) of 4.4 are required for the following reasons.

- There are two distinct elements in the scope of the plan. The first identifies the medical device, the other identifies the phase of the life cycle for which each element of the plan is applicable. By defining the scope, the manufacturer establishes the baseline on which all the risk management activities are built.

- Allocation of responsibilities and authorities is needed to ensure that no responsibility is omitted.

- Review of activities such as risk management is included as a generally recognised responsibility of management.

- The criteria for risk acceptability are fundamental to risk management and should be decided upon before risk analysis begins. This helps to make the risk evaluation in Clause 6 objective.

- After implementing all risk control measures, the manufacturer is required to evaluate the overall impact of all residual risks together. The evaluation method and the criteria for acceptability of the overall residual risk should be decided upon before this evaluation is performed. This helps to make the evaluation of overall residual risk in Clause 8 objective.

- Verification is an essential activity and is required by 7.2. Planning this activity helps to ensure that essential resources are available when required. If verification is not planned, important parts of the verification could be neglected.

- Methods for the collection and review of production and post-production information need to be established so that there is a formal and appropriate way to feed back production and postproduction information into the risk management process.

The requirement to keep a record of changes is to facilitate audit and review of the risk management process for a particular medical device.

###### A.2.4.5 	Risk management file

This document uses this term to signify where the manufacturer can locate or find the locations of all the records and other documents applicable to risk management. This facilitates the risk management process and enables more efficient auditing to this document. Traceability is necessary to demonstrate that the risk management process has been applied to each identified hazard.

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Completeness is very important in risk management. An incomplete task can mean that an identified hazard is not controlled and harm can be the consequence. The problem can result from incompleteness at any step in risk management, e.g. unidentified hazards, risks not assessed, unspecified risk control measures, risk control measures not implemented, or risk control measures that prove ineffective. Traceability is needed to ensure completeness of the risk management process.

##### A.2.5 Risk analysis

###### A.2.5.1 	Risk analysis process

Note 1 of 5.1 describes how to deal with the availability of a risk analysis for a similar medical device. When adequate information already exists, this information can be applied to save time, effort and resources. Users of this document need to be careful, however, to assess systematically the previous work for applicability to the current risk analysis.

The details required by a), b), and c) form the basic minimum data set for ensuring traceability and are important for management reviews and for subsequent audits. The requirement in c) also helps clarify what is in the scope of the analysis and verifies completeness.

###### A.2.5.2 	Intended use and reasonably foreseeable misuse

The intended use of the medical device is an important aspect and is the starting point of the risk analysis. This should include the elements listed in the note to 3.6, where appropriate. The manufacturer should also consider the intended user(s) of the medical device, e.g., whether a lay user or a trained medical professional will use the medical device. This analysis should consider that medical devices can also be used in situations other than those intended by the manufacturer and in situations other than those foreseen when the idea for a medical device was first conceived. It is important that the manufacturer tries to look into the future to see the hazards due to potential uses of their medical device and also the reasonably foreseeable misuse.

###### A.2.5.3	Identification	of	characteristics	related	to	safety

This step forces the manufacturer to think about all the characteristics that could affect the safety of the medical device. These characteristics can be qualitative or quantitative and can be related to the operating principle of the medical device, its intended use and/or the reasonably foreseeable misuse. Such characteristics can relate to the performance or operating principle of the medical device, the measuring function or the sterility of the medical device, the materials used for parts coming into contact with the patient, the use of radiation for diagnostic or therapeutic purposes, or other. Where applicable, the limits of those characteristics need to be considered as well, because the operation and/or safety of the medical device could be affected when those limits are exceeded.

###### A.2.5.4	Identification	of	hazards and hazardous situations

This step requires the manufacturer to be systematic in the identification of anticipated hazards in both normal and fault conditions. The identification should be based upon the intended use and reasonably foreseeable misuse identified in 5.2 and the characteristics related to safety identified in 5.3.

A risk can only be assessed and managed once a hazardous situation has been identified. Documenting the reasonably foreseeable sequences of events that can transform a hazard into a hazardous situation allows this to be done systematically. Annex C aims to assist manufacturers in identifying hazards and hazardous situations. Typical hazards are listed and the relationships between hazards, foreseeable sequences of events, hazardous situations and associated possible harm are demonstrated.

A.2.5.5 	Risk estimation

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

This is the final step of risk analysis. The difficulty of this step is that risk estimation is different for every hazardous situation that is under investigation as well as for every medical device. Therefore, this subclause was written generically. Because hazards can occur both when the medical device functions normally and when it malfunctions, one should look closely at both situations. In practice, both components of risk, probability of occurrence and severity of harm, should be analysed separately. When a manufacturer uses a systematic way of categorizing the severity levels or the probability of occurrence of harm, the categorization scheme should be defined and recorded in the risk management file. This enables the manufacturer to treat equivalent risks consistently and serves as evidence that the manufacturer has done so.

Some hazardous situations occur because of systematic faults or sequences of events. There is no consensus on how to calculate the probability of a systematic fault. Where the probability of occurrence of harm cannot be calculated, hazards still have to be addressed and listing resulting hazardous situations separately allows the manufacturer to focus on reducing the risks due to these hazardous situations.

Frequently, good quantitative data are not readily available, especially in development of an entirely new medical device or for security risks. The suggestion that risk estimation should be done only in a quantitative way has therefore been avoided.

##### A.2.6 Risk evaluation

Decisions have to be made about the acceptability of risk. Manufacturers can use the estimated risks and evaluate them using the criteria for risk acceptability defined in the risk management plan. They can investigate the risks to determine which ones need to be controlled. Clause 6 was carefully worded to allow the user of this document to avoid unnecessary work.

##### A.2.7 Risk control

###### A.2.7.1 	Risk control option analysis

Often there will be more than one way to reduce a risk. There are three mechanisms listed, which are all standard risk reduction measures and are derived from ISO/IEC Guide 63:2019[2]. The priority order listed is important. This principle is found in several places, including IEC TR 60513[11] and local or regional regulations. Inherently safe design and manufacture is the first and most important option in the risk control option analysis, because design solutions inherent to the characteristics of the medical device are likely to remain effective, whereas experience has shown that even well-designed guards and protective measures can fail or be violated and information for safety might not be followed. If practicable, the medical device should be designed and manufactured to be inherently safe. If this is not practicable, then protective measures such as barriers or alarms are appropriate. The third option is to provide information for safety such as a written warning or contra-indication. Training to users can be an important aspect of delivering information for safety. The manufacturer can consider to provide mandatory training for the intended users.

The manufacturing process can contribute to risks, for example originating from contamination of components, residues of hazardous substances used in the process, or mix-up of parts. Such risks can be controlled by designing the manufacturing process to be inherently safe (e.g. eliminating hazardous substances or using separate production lines) or by applying protective measures (e.g. visual inspection steps in the process).

It is recognised that one possible result of the risk control option analysis could be that there is no practicable way of reducing the risk to acceptable levels according to the pre-established criteria for risk acceptability. For example, it could be impractical to design a life-supporting medical device with such an acceptable residual risk. In this case, a benefit-risk analysis can be carried out as described in 7.4 to determine whether the benefit of the medical device, to the patient, outweighs the residual risk. This option is included at this point in the document to make sure that every effort was first made to reduce risks to the pre-established acceptable levels.

###### A.2.7.2 	Implementation of risk control measures

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Two distinct verifications are included. The first verification is required to make sure that the risk control measure has been implemented in the final design of the medical device or in the manufacturing process. The second verification is required to ensure that the risk control measure (including information for safety) as implemented actually reduces the risk. In some instances, a validation study can be used for verifying the effectiveness of the risk control measure.

Obtaining sufficient data and information for risk estimation can be difficult, resulting in uncertainty of the residual risk evaluation. It can therefore be practical for the manufacturer to focus effort on verification of effectiveness of risk control measures to establish a convincing residual risk evaluation. Level of effort should be commensurate with the level of risk. Testing with users might be needed to verify the effectiveness of the risk controls, for example usability testing (see IEC 62366-1[3]), clinical investigation of medical devices (see ISO 141551)[6]) or clinical performance studies for in vitro diagnostic medical devices (see ISO 20916[8]). A usability test can verify effectiveness of information for safety and a test according to a test standard can verify effectiveness of designed risk control measures related to, for example, mechanical strength.

###### A.2.7.3 	Residual risk evaluation

A check was introduced here to determine whether the implemented risk control measures have made the risk acceptable. If the risk exceeds the acceptability criteria established in the risk management plan, the manufacturer is instructed to investigate additional risk control measures. This iterative procedure should be continued until further risk control is not practicable and the residual risk does not exceed the acceptability criteria established in the risk management plan.

###### A.2.7.4 	Benefit-risk analysis

There can be particular hazardous situations for which the risk exceeds the manufacturer’s criteria for risk acceptability. This subclause enables the manufacturer to provide a high-risk medical device for which they have done a careful evaluation and can show that the benefit of the medical device outweighs the risk. However, this subclause cannot be used to weigh residual risks against economic advantages or business advantages (i.e. for business decision making).

###### A.2.7.5 	Risks arising from risk control measures

This subclause recognises that risk control measures alone or in combination might introduce a new and sometimes quite different hazard, and that risk control measures introduced to reduce one risk might increase another risk.

###### A.2.7.6 	Completeness of risk control

At this point, the risks of all the hazardous situations should have been evaluated. This check was introduced to ensure that no hazardous situations were left out in the intricacies of a complex risk analysis.

#### A.2.8 Evaluation of overall residual risk

During the process defined by Clauses 5 to 7, manufacturers identify hazards and hazardous situations, evaluate the risks, and implement risk control measures in their medical device design one at a time. This is the point where the manufacturer has to step back, consider the combined impact of all individual residual risks, and make a decision as to whether to proceed with the medical device. It is possible that the overall residual risk exceeds the manufacturer’s criteria for risk acceptability, even though individual residual risks do not. This is particularly true for complex systems and medical devices with a large number of risks. The method to evaluate the overall residual risk as defined in the risk management plan includes balancing the overall residual risk against the benefits of the medical device. This is particularly relevant in determining whether a high-risk, but highly beneficial, medical device should be marketed.

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

The manufacturer is responsible for providing users with relevant information on significant residual risks, so that they can make informed decisions on the use of the medical device. Thus, manufacturers are instructed to include pertinent information on residual risks in the accompanying documentation. However, it is the manufacturer’s decision as to what and how much information should be provided. This requirement is consistent with the approach taken in many countries and regions.

##### A.2.9 Risk management review

The risk management review is an important step before the commercial release of the medical device. The final results of the risk management process, as obtained by executing the risk management plan, are reviewed. The risk management report contains the results of this review and is a crucial part of the risk management file. The report serves as the high-level document that provides evidence that the manufacturer has ensured that the risk management plan has been satisfactorily fulfilled and the results confirm that the required objective has been achieved. Subsequent reviews of the execution of the risk management plan and updates of the risk management report can be needed during the life cycle of the medical device, as a result of the execution of production and post-production activities.

#### A.2.10    Production and post-production activities

It cannot be emphasized too often that risk management does not stop when a medical device goes into production. Risk management often begins with an idea, before there is any physical manifestation of the medical device. Manufacturers collect information from many sources, including experience with similar medical devices and technologies. Risk estimation is refined throughout the design process and can be made more accurate when a functioning prototype is built. However, no amount of modelling can substitute for an actual medical device in the hands of actual users.

Therefore, the manufacturer needs to collect and review production and post-production information and evaluate its relevance to safety. The information can relate to new hazards or hazardous situations, and/or can affect their risk estimates or the balance between benefit and overall residual risk. Either can impact the manufacturer’s risk management decisions. The manufacturer should also take into account considerations of the generally acknowledged state of the art, including new or revised standards. When the information is determined to be relevant to safety, the risk management process requires that it be considered as an input for modification of the medical device, and also as an input to improve the process itself. With effective production and post-production activities, the risk management process truly becomes an iterative closed-loop process to ensure the continued safety of the medical device.

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

In reply to feedback and requests for additional guidance and in response to changing regulatory requirements, the requirements for production and post-production activities are elaborated in more detail in this third edition. The clause is divided into subclauses. More sources of information are listed, including information on the generally acknowledged state of the art and feedback from the supply chain. The latter includes suppliers of components or subsystems, and also third-party software. The possible need for actions regarding medical devices already on the market is made more explicit. The conditions under which follow-up actions need to be considered, are extended with changes in the state of the art that can be relevant to safety, such as alternative medical devices and/or therapies becoming available on the market, as well as changes in risk perception or risk acceptability.

## Annex B

(informative)

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Risk management process for medical devices

### B.1 Correspondence between second and third editions

The numbering of clauses and subclauses has changed with this third edition of ISO 14971. Table B.1 provides the correspondence between clauses and subclauses in the second edition ISO 14971:2007 and those in the third edition ISO 14971:2019. This table is provided to assist users of this document in transitioning from the second to the third edition and to facilitate updating of references to ISO 14971 in other documents.

Table B.1 — Correspondence between elements of ISO 14971:2007 and ISO 14971:2019

| ISO 14971:2007 | ISO 14971:2019 |
| --- | --- |
| Introduction | Introduction |
| 1 Scope | 1 Scope |
| (New clause) | 2 Normative references |
| 2 Terms and definitions | 3 Terms and definitions |
| 2.1 accompanying document | 3.1 accompanying documentation |
| (New definition) | 3.2 benefit |
| 2.2 harm | 3.3 harm |
| 2.3 hazard | 3.4 hazard |
| 2.4 hazardous situation | 3.5 hazardous situation |
| 2.5 intended use intended purpose | 3.6 intended use intended purpose |
| 2.6 in vitro diagnostic medical device IVD medical device | 3.7 in vitro diagnostic medical device IVD medical device |
| 2.7 life-cycle | 3.8 life cycle |
| 2.8 manufacturer | 3.9 manufacturer |
| 2.9 medical device | 3.10 medical device |
| 2.10 objective evidence | 3.11 objective evidence |
| 2.11 post-production | 3.12 post-production |
| 2.12 procedure | 3.13 procedure |
| 2.13 process | 3.14 process |
| (New definition) | 3.15 reasonably foreseeable misuse |
| 2.14 record | 3.16 record |
| 2.15 residual risk | 3.17 residual risk |

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Table B.1 (continued)

| ISO 14971:2007 | ISO 14971:2019 |
| --- | --- |
| 2.16 risk | 3.18 risk |
| 2.17 risk analysis | 3.19 risk analysis |
| 2.18 risk assessment | 3.20 risk assessment |
| 2.19 risk control | 3.21 risk control |
| 2.20 risk estimation | 3.22 risk estimation |
| 2.21 risk evaluation | 3.23 risk evaluation |
| 2.22 risk management | 3.24 risk management |
| 2.23 risk management file | 3.25 risk management file |
| 2.24 safety | 3.26 safety |
| 2.25 severity | 3.27 severity |
| (New definition) | 3.28 state of the art |
| 2.26 top management | 3.29 top management |
| 2.27 use error | 3.30 use error |
| 2.28 verification | 3.31 verification |
| 3 General requirements for risk management | 4 General requirements for risk management system |
| 3.1 Risk management process | 4.1 Risk management process |
| 3.2 Management responsibilities | 4.2 Management responsibilities |
| 3.3 Qualification of personnel | 4.3 Competence of personnel |
| 3.4 Risk management plan | 4.4 Risk management plan |
| 3.5 Risk management file | 4.5 Risk management file |
| 4 Risk analysis | 5 Risk analysis |
| 4.1 Risk analysis process | 5.1 Risk analysis process |
| 4.2 Intended use and identification of characteristics related to the safety of the medical device | 5.2 Intended use and reasonably foreseeable misuse |
|  | 5.3 Identification of characteristics related to safety |
| 4.3 Identification of hazards | 5.4 Identification of hazards and hazardous situations |
| 4.4 Estimation of the risk(s) for each hazardous situation | 5.5 Risk estimation |
| 5 Risk evaluation | 6 Risk evaluation |
| 6 Risk control | 7 Risk control |
| 6.1 Risk reduction | (Subclause deleted) |
| 6.2 Risk control option analysis | 7.1 Risk control option analysis |
| 6.3 Implementation of risk control measure(s) | 7.2 Implementation of risk control measures |
| 6.4 Residual risk evaluation | 7.3 Residual risk evaluation |
| 6.5 Risk/benefit analysis | 7.4 Benefit-risk analysis |
| 6.6 Risks arising from risk control measures | 7.5 Risks arising from risk control measures |
| 6.7 Completeness of risk control | 7.6 Completeness of risk control |
| 7 Evaluation of overall residual risk acceptability | 8 Evaluation of overall residual risk |
| 8 Risk management report | 9 Risk management review |

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Table B.1 (continued)

| ISO 14971:2007 | ISO 14971:2019 |
| --- | --- |
| 9 Production and post-production information | 10 Production and post-production activities |
|  | 10.1 General |
|  | 10.2 Information collection |
|  | 10.3 Information review |
|  | 10.4 Actions |
| Annex A Rationale for requirements | Annex A Rationale for requirements |
| Annex B Overview of the risk management process for medical devices | Annex B Risk management process for medical devices |
| Annex C Questions that can be used to identify medical device characteristics that could impact on safety | Moved to ISO/TR 24971 |
| Annex D Risk concepts applied to medical devices |  |
| Annex E Examples of hazards, foreseeable sequences of events and hazardous situations | Annex C Fundamental risk concepts |
| Annex F Risk management plan | Moved to ISO/TR 24971 |
| Annex G Information on risk management techniques |  |
| Annex H Guidance on risk management for in vitro diagnostic medical devices |  |
| Annex I Guidance on risk analysis process for biological hazards | (Annex deleted) |
| Annex J Information for safety and information about residual risk | Moved to ISO/TR 24971 |
| Bibliography | Bibliography |

#### B.2 Risk management process overview

Figure B.1 is provided to give the user of this document an overview of the risk management process. It is for illustrative purposes only. As indicated in Figure B.1, the process needs to be iterative, covering each risk in turn, and returning to earlier steps if risk control measures introduce new hazards or hazardous situations, or if new information becomes available.

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Figure B.1 — Overview of risk management activities as applied to medical devices

Annex C

(informative)

## Fundamental risk concepts

### C.1 General

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

This document requires the manufacturer to compile a list of known and foreseeable hazards associated with the medical device in both normal and fault conditions and to consider the foreseeable sequences of events that can produce hazardous situations and harm. According to the definitions, a hazard cannot result in harm until such time as a sequence of events or other circumstances (including normal use) lead to a hazardous situation. At this point, the risk can be assessed by estimating both severity and probability of occurrence of harm that could result (see Figure C.1). The probability of occurrence of harm can be expressed as a combination of separate probabilities (P1, P2) or as a single probability (P). A decomposition into P1 and P2 is not mandatory.

NOTE 1   Depending on the complexity of the medical device, a hazard can lead to multiple hazardous situations, and each hazardous situation can lead to multiple harms.

NOTE 2   The probability of occurrence of harm (P) can be composed of separate P1 and P2 values.

NOTE 3   The thin arrows represent elements of risk analysis and the thick arrows depict how a hazard can lead to harm

Figure C.1 — Pictorial example of the relationship between hazard, sequence of events, hazardous situation and harm (from ISO/IEC Guide 63:2019[2])

A good starting point for this compilation is a review of experience with the same and similar types of medical devices. The review should take into account a manufacturer’s own experience and, where appropriate, the experience of other manufacturers as reported in adverse event databases, publications, scientific literature and other available sources. This type of review is particularly useful for the identification and listing of typical hazards and hazardous situations for a medical device and the associated harm that can occur. Next, this listing and aids such as the list of examples in Table C.1 can be used to compile an initial list of hazards.

It is then possible to begin identification of some of the sequences of events that together with hazards could result in hazardous situations and harm. Since many hazards might never result in harm and can be eliminated from further consideration, it could be useful to perform this analysis by starting with the harm that the medical device might cause and work backwards to the hazardous situations, hazards and initiating causes. However, although this approach is useful for the reason described, it should be recognised that it is not a thorough analysis. Many sequences of events will only be identified by the systematic use of risk analysis techniques (such as those described in ISO/TR 24971[9]). Analysis and identification are further complicated by the many events and circumstances that have to be taken into consideration such as those listed in Table C.2. Thus, more than one risk analysis technique, and especially complementary techniques, are often used to complete a comprehensive analysis. Table C.3 provides examples of the relationship between hazards, sequences of events, hazardous situations, and harm.

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Although compilation of the lists of hazards, hazardous situations and sequences of events should be completed as early as possible in the design and development process to facilitate risk control, in practice identification and compilation is an ongoing activity that continues throughout the life cycle of the medical device through post-production to disposal.

This annex provides a non-exhaustive list of possible hazards that can be associated with different medical devices (Table C.1) and a list of events and circumstances (Table C.2) that can result in hazardous situations, which can result in harm. Table C.3 provides examples in a logical progression of how a hazard can be transformed into a hazardous situation and produce harm by a sequence of events or circumstances.

Recognising how hazards progress to hazardous situations is critical for estimating the probability of occurrence and severity of harm that could result. An objective of the process is to compile a comprehensive set of hazardous situations. The identification of hazards and sequences of events are stepping stones to achieve this. The lists in the tables in this annex can be used to aid in the identification of hazardous situations. What is called a hazard needs to be determined by the manufacturer to suit the particular analysis.

### C.2 Examples of hazards

The list in Table C.1 can be used to assist in the identification of hazards associated with a particular medical device, which could ultimately result in harm.

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Table C.1 — Examples of hazards

| Energy hazards | Biological and chemical hazards | Performance-related hazards |
| --- | --- | --- |
| Acoustic energy — infrasound — sound pressure — ultrasonic Electric energy Electric fields Leakage current — earth leakage — enclosure leakage Magnetic fields Static discharge Voltage Mechanical energy Kinetic energy — falling objects — high pressure fluid injection — moving parts — vibrating parts Potential (stored) energy — bending — compression — cutting, shearing — gravitational pull — suspended mass — tension — torsion Radiation energy Ionizing radiation — accelerated particles (alpha particles, electrons, protons, neutrons) — gamma — x-ray Non-ionizing radiation — infrared — laser — microwave — ultraviolet Thermal energy Cryogenic effects Hyperthermic effects | Biological agents Bacteria Fungi Parasites Prions Toxins Viruses Chemical agents Carcinogenic, mutagenic, reproductive Caustic, corrosive — acidic — alkaline — oxidants Flammable, combustible, explosive Fumes, vapors Osmotic Particles (including micro- and nanoparticles) Pyrogenic Solvents Toxic — asbestos — heavy metals — inorganic toxicants — organic toxicants — silica Immunological agents Allergenic — antiseptic substances — latex Immunosuppressive Irritants — cleaning residues Sensitizing | Data — access — availability — confidentiality — transfer — integrity Delivery — quantity — rate Diagnostic information — examination result — image artefacts — image orientation — image resolution — patient identity / information Functionality — alarm — critical performance — measurement |

### C.3 Examples of events and circumstances

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

In order to identify foreseeable sequences of events, it is often useful to consider events and circumstances that can cause them. Table C.2 provides examples of events and circumstances, organized into general categories. Although the list is certainly not exhaustive, it is intended to demonstrate the many different types of events and circumstances that need to be taken into account to identify the foreseeable sequences of events for a medical device.

Table C.2 — Examples of events and circumstances

| General category | Events and circumstances |
| --- | --- |
| Requirements | Inadequate specification of: — design parameters — operating parameters — performance requirements — in-service requirements (e.g. maintenance, reprocessing) — end of life |
| Manufacturing processes | Insufficient control of: — manufacturing processes — changes to manufacturing processes — materials — materials compatibility information — subcontractors |
| Transport and storage | Inadequate packaging Contamination or deterioration Inappropriate environmental conditions |
| Environmental factors | Physical factors (e.g. heat, pressure, time) Chemical factors (e.g. corrosion, degradation, contamination) Electromagnetic fields (e.g. susceptibility to electromagnetic disturbance) Inadequate supply of power Inadequate supply of coolant |
| Cleaning, disinfection and sterilization | Lack of validated procedures Inadequate specification of requirements Inadequate performance of cleaning, disinfection or sterilization |
| Disposal and scrapping | No or inadequate information provided Use error |
| Formulation | Biodegradation Biocompatibility No information or inadequate specification provided Incorrect formulations Use error |

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Table C.2 (continued)

| General category | Events and circumstances |
| --- | --- |
| Usability | Confusing or missing instructions for use Complex or confusing control system Ambiguous or unclear state of the medical device Ambiguous or unclear presentation of settings, measurements or other information Misrepresentation of results Insufficient visibility, audibility or tactility Poor mapping of controls to actions, or of displayed information to actual state Controversial modes or mapping as compared to existing equipment Use by unskilled or untrained personnel Insufficient warning of side effects Inadequate warning of hazards associated with re-use of single-use medical devices Incorrect measurement and other metrological aspects Incompatibility with consumables, accessories, other medical devices Incorrect patient identification Slips, lapses and mistakes |
| Functionality | Loss of electrical or mechanical integrity Deterioration in performance (e.g. gradual occlusion of fluid or gas path, change in resistance to flow, electrical conductivity) as result of ageing, wear and repeated use Failure of a component due to ageing, wear or fatigue |
| Security | Unsecured data ports that are externally accessible (e.g. network, serial or USB ports) Data without encryption Software vulnerabilities that can be exploited Software updates without authenticity confirmation |

### C.4 Examples of relationships between hazards, foreseeable sequences of events, hazardous situations and the harm that can occur

Table C.3 illustrates the relationship between hazards, foreseeable sequences of events, hazardous situations and harm for some simplified examples. Remember that one hazard can result in more than one harm and that more than one sequence of events can give rise to a hazardous situation.

The decision on what constitutes a hazardous situation needs to be made to suit the particular analysis being carried out. In some circumstances it can be useful to describe a cover being left off a high voltage terminal as a hazardous situation, in other circumstances the hazardous situation can be more usefully described as when a person is in contact with the high voltage terminal.

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Table C.3 — Relationship between hazards, foreseeable sequences of events, hazardous situations and the harm that can occur

| Hazard |  | Foreseeable sequence of events | Hazardous situation | Harm |
| --- | --- | --- | --- | --- |
| Electromagnetic energy (high voltage) | (1) | Electrode cable unintentionally plugged into power line receptacle | Line voltage appears on electrodes | Serious burns Heart fibrillation |
| Chemical (volatile solvent, embolus) | (1) (2) | Incomplete removal of volatile solvent used in manufacturing Solvent residue converts to gas at body temperature | Development of gas embolism (bubbles in the blood stream) during dialysis | Infarct Brain damage |
| Biological (microbial contamination) | (1) (2) | Inadequate instructions provided for decontaminating re-used anaesthesia tubing Contaminated tubing used during anaesthesia | Bacteria released into airway of patient during anaesthesia | Bacterial infection |
| Functionality (no delivery) | (1) (2) | Electrostatically charged patient touches infusion pump Electrostatic discharge (ESD) causes pump and pump alarms to fail | Failure to deliver insulin to patient with elevated blood glucose level, no warning given | Minor organ damage Decreased consciousness |
| Functionality (no output) | (1) (2) | Implantable defibrillator battery reaches the end of its useful life Inappropriately long interval between clinical follow-up visits | Defibrillator cannot deliver shock when an arrhythmia occurs | Death |
| Measurement (incorrect information) | (1) (2) | Measurement error No detection by user | Incorrect information reported to clinician, leading to misdiagnosis and/ or lack of proper therapy | Progression of disease Serious injury |

## Bibliography

- ISO/IEC Guide 51:2014, Safety aspects — Guidelines for their inclusion in standards

- ISO/IEC Guide 63:2019, Guide to the development and inclusion of aspects of safety in international standards for medical devices

- ISO 9000:2015, Quality management systems — Fundamentals and vocabulary

- ISO 10993-1, Biological evaluation of medical devices — Part 1: Evaluation and testing within a risk management process

- ISO 13485:2016, Medical devices — Quality management systems — Requirements for regulatory purposes

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

ISO 14155, Clinical investigation of medical devices for human subjects — Good clinical practice

- ISO 18113-1:2009, In vitro diagnostic medical devices — Information supplied by the manufacturer (labelling) — Part 1: Terms, definitions and general requirements

- ISO 20916, In vitro diagnostic medical devices — Clinical performance studies using specimens from human subjects — Good study practice

- ISO/TR 24971, Medical devices — Guidance on the application of ISO 14971

- ISO 31000, Risk management — Guidelines

- IEC/TR 60513, Fundamental aspects of safety standards for medical electrical equipment

- IEC 60601-1, Medical electrical equipment — Part 1: General requirements for basic safety and essential performance

- IEC 62366-1:2015, Medical devices — Part 1: Application of usability engineering to medical devices

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

[Blank page] [Blank page]

Internal network licence for 2 users. Order No. 438356, 9.03.2021.

Right of Use of this electronic document: Noncomplicity AB, the right to reproduce and distribute belongs to the Estonian Centre for Standardisation and Accreditation.

EVS, Internal network licence for 2 users. Order No. 438356, 9.03.2021.

standard@evs.ee www.evs.ee

)  Under preparation. Stage at the time of publication ISO/FDIS 14155:2019.